## Slide 1
Development of a Liquid-Level Monitoring System and Molecular Oxygen Sensor
Stephanie Muggler, Leland Bruce Gee Jr.
LCLS, SLAC National Accelerator Laboratory, Menlo Park, California 94025
Sensor Considerations
In order to fit into the liquid jet enclosure, the monitor must fit into both a 50 mL and 15 mL Falcon tube.
The sensor must also be able to communicate with its user via python.
Our solution was to use a non-contact sensor that is programmable via Arduino and Pyserial.
Why Liquid Level Monitoring?
Volatile solvents evaporate easily in the LCLS liquid jet enclosure, resulting in constant concentration chances as well as adding downtime to experiments. This creates uncertainty in LCLS experiments, creating a need for a way to monitor liquid levels within these reservoirs, hence the need for a liquid level monitor.
Molecular Oxygen Sensing at LCLS
Due to many inorganic and bioinorganic systems being intolerant to oxygen environments and no currently no way to assess how anaerobic the liquid jet system is. By performing benchtop anaerobic UV-Vis experiments “offline” we can calibrate the sensitivity of potential oxygen sensors to ensure anaerobicity. This experiment focused on calibrating Methyl Viologen as a possible oxygen sensor.
Building the Monitor
= 1.1”
= 8 mL
(50)
= 2 mL
(15)
Since the sensor is a binary capacitance sensor and can’t display liquid level as percentages, it was best to position it to only alert when the liquid level fell below a certain amount. This would mean an 8 mL drop for 50 mL, and 2 mL for 15 mL.
The Arduino monitor was then hooked up to BNC cables and placed in a 3D printed case.
Designing the Experiment
Despite sparging separately with both N2 and Ar, the methyl viologen solution still lost color, due to oxidation, after a few minutes. This indicates that the crimp bottles weren’t completely airtight.
Anaerobic Solution immediately after dithionite is added
Solution around 4 minutes after
The solution was to place the UV-Vis fiber optics and cuvette holder in a glove bag, and conduct the experiment inside.
Glove Bag with UV-Vis fiber-optic cables
UV-Vis
Spectrometer
O2 / Moisture
Removal
Column
[MV]+ ⇌ [MV]+2 + e-
Deep Blue
Clear
Analyzing UV-Vis Data
Change in Methyl Viologen Absorption at Different Wavelengths
A change in absorption was shown most prominently at 285, 395, and 600 nm; graphing these changes as a function of percent oxygen gives a linear trend, as predicted by Beer’s Law.
There are some discrepancies, as at 600 nm, the trend was more curved than linear, most likely due to not adding the exact amount of oxygen.
Future Improvements
A future goal would be to connect the Liquid Level monitor to the liquid jet or to regenerate reduced Methyl Viologen using light with EDTA and proflavine.
Acknowledgements
Thank you to both Leland and SLAC LCLS for hosting and mentoring me this summer. Thank you so much!
O2
Na2S2O4
600 nm